## 3D Printed Parts
No | Name | Description | Amount
1 | Centrifuge Box - lv01 (Bottom) v.1 | Motor & Arduino Holder (Vibration damping pad inside motor holder is recommended) | 1
2 | Centrifuge Box - lv01 (Bottom) v.2 | Motor & Arduino Holder [MIRRORED] (Only Used if v.1 is not used) | 0-1
3 | Centrifuge Box - lv02 | Frame Wall to hold Buttons and Display | 1
4 | Centrifuge Box - lv03 | Frame Wall (Plain Simple) | 1
5 | Centrifuge Box - lv04 (top) | Frame Wall (Plain Simple) | 1
6 | Centrifuge Box - lv03 Ball Bearing House Cross | To hold the One Way Clutch Ball Bearing | 1
7 | Centrifuge Box - lv04 Ball Bearing House Cross | To hold the lower Ball Bearing stabilizer | 1
8 | Centrifuge Box - Top Ball Bearing House Cross v.1 | To hold the Upper Ball Bearing stabilizer (No hall sensor support) (Only Used if v.2 is not used) | 0-1
9 | Centrifuge Box - Top Ball Bearing House Cross v.2 | To hold the Upper Ball Bearing stabilizer (KS0020 hall sensor support) | 1
10 | Centifuge Test Tube Holder - No Holes | Template to use if custom tube sizes are needed | 0
11 | Centifuge Test Tube Holder - With Holes | 4 Tube Holes for d:29mm, l: 116mm tubes | 1
12 | Centrifuge Box - Lid | Top Cover to protect user if something goes wrong | 1
13 | Support Piece for Digit Display and Btn Board | Helper parts to hold and support the button and display parts ( (needs to be glued)) | 1
14 | Box Corner Stand | To be attached to metal frame and fit the Centrifuge box between | 4
15 | TPU Rubber Feet | Controls the natural shaking and movement of the centrifuge | 4
16 | Feet Extension | Gives more stability to the centrifuge (Adding Anti-slip tape is recommended) | 4

## Electronics
No | Name | Description | Amount
1 | 4-Digit 7-Segment Display | Displays 4 digit numbers (0000-9999) or text for Speed state and current RPM | 1
2 | Breadboard (Normal Size) | Shared VCC and GND Central Board | 1
3 | Breadboard (Small Size) | Button Holder | 1
4 | BTS7960 motor driver | Motor Driver Circuit Board to control Power distribution to the DC-Motor | 1
5 | Buttons (ON/OFF) and (RPM STEP) | Classic 4-Pin Buttons for controlling Speed state and On/Off | 2
6 | DC Motor 775 10000RPM 12V/24V 5-10A | A powerful High Speed DC-Motor (15000-20000 RPM at No-Load) | 1
7 | Hall Sensor (KS0020) | Sensor that reacts to passing magnetic pulses in order to count RPM | 1
8 | Low Pass Filter Capacitor | Used to make the Hall sensor less sensitive to DC-Motor magnetic field (Sensor Pin to GND) | 1
9 | Arduino UNO R3 PCB | The main controller of all electronic parts, running the firmware | 1

## Mechanical Hardware
No | Name | Description | Amount
1 | 2020 Extrusion (300mm) | To use as Centrifuge frame for stabilizerand centrifuge when running | 4
2 | Metal Frame Corner Connector | To connect the 4 2020 extrusion bars | 4
3 | Bolts M3 x 8 | To fasten the 4 Metal frame corner connector to the 2020 extrusions | 20
4 | 20-Series Square Nut M3 | To fasten the 4 Metal frame corner connector to the 2020 extrusions | 20
5 | Bolts M4 x 8 | To fasten the 4 Box Corner Stands to the 2020 extrusions | 8
6 | 20-Series Square Nut M3 | To fasten the 4 Box Corner Stands to the 2020 extrusions | 8
7 | Bolts M4 x 8 | To secure the Flange Hub to the Centifuge Test Tube Holder | 4
8 | Nut M4 | To secure the Flange Hub to the Centifuge Test Tube Holder | 4
9 | Flange Shaft Fitting Metal Hub | To secure the Centifuge Test Tube Holder to the Shaft | 1
10 | Flange Secure Bolt | To secure the Centifuge Test Tube Holder to the Shaft | 2
11 | Bolts M2.5 x 6 | To fasten the Hull Sensor in place below (1.5mm) the spinning magnet | 1
12 | Nut M2.5 | To fasten the Hull Sensor in place below (1.5mm) the spinning magnet | 1
13 | Magnet (Tiny) | Fastened under the Test Tube Holder to allow the Hall sensor count the RPM | 2
14 | Nut M2 | To counter weight the magnets on the Test Tube Holder | ~2
15 | Shaft | The spinning shaft connected to the motor and the Test tube Holder | 1
16 | Coupling M5 to M8 | To connect the motor to the shaft | 1
17 | One Way Bearing Clutch (id:8mm;od:22mm;w:7mm) | Connected closest to the motor to allow for free spinning when motor stops | 1
18 | Ball Bearing (id:8mm;od:22mm;w:7mm) | Lower and Upper Ball bearing for stabilization of the shaft | 2

## Wires
No | Name | Description* | Amount
1 | Wire – Motor Supply + | External Power Supply + → IBT-2 B+ | 1
2 | Wire – Motor Supply − | External Power Supply − → IBT-2 B− | 1
3 | Wire – Common GND Bus | Arduino GND → Common Ground Distribution | 1
4 | Wire – IBT-2 Logic VCC | Arduino 5V → IBT-2 VCC | 1
5 | Wire – IBT-2 Logic GND | Arduino GND → IBT-2 GND | 1
6 | Wire – IBT-2 RPWM | Arduino D5 → IBT-2 RPWM | 1
7 | Wire – IBT-2 LPWM | Arduino D6 → IBT-2 LPWM | 1
8 | Wire – IBT-2 R-EN | Arduino D7 → IBT-2 R-EN | 1
9 | Wire – IBT-2 L-EN | Arduino D8 → IBT-2 L-EN | 1
10 | Wire – Motor + | IBT-2 M+ → Motor + | 1
11 | Wire – Motor − | IBT-2 M− → Motor − | 1
12 | Wire – Hall VCC | Arduino 5V → Hall Sensor VCC | 1
13 | Wire – Hall GND | Arduino GND → Hall Sensor GND | 1
14 | Wire – Hall Signal | Hall Sensor Signal → Arduino D2 | 1
15 | Capacitor – 100nF Ceramic | Hall Signal → GND (Noise suppression capacitor) | 1
16 | Wire – Display VCC | Arduino 5V → TM1637 VCC | 1
17 | Wire – Display GND | Arduino GND → TM1637 GND | 1
18 | Wire – Display CLK | Arduino D3 → TM1637 CLK | 1
19 | Wire – Display DIO | Arduino D4 → TM1637 DIO | 1
20 | Wire – BTN1 Signal | BTN1 → Arduino Digital Input | 1
21 | Wire – BTN1 GND | BTN1 → Arduino GND | 1
22 | Wire – BTN2 Signal | BTN2 → Arduino Digital Input | 1
23 | Wire – BTN2 GND | BTN2 → Arduino GND | 1
* If firmware pin assignments are modified, this wiring list must be updated accordingly.